## Slide 1: Rinstrum Verification Officers Refresher Training 2022
X00A-837-201

## Slide 2: Content – RVO Refresher Training 2022
1. Introduction
  1.1 Guidelines and Standards
  1.2 EU Directives
  1.3 Prerequisites
  1.4 Rinstrum Manufacture Conformity Assessment Procedure
  1.5 Rinstrum MCA Team
2. MCA Procedures and Instructions
  3.1 RVO MCA Procedure
  3.2 Check List for the Execution of MCA (X00A-808)
  3.3 Comp Calc X00A-805
  3.4 Verification Tests X00A-804
  3.5 Seals and Labels
  3.6 Documents to Rinstrum
  3.7 Declaration
  3.8 Clarification and Rejection
3. Rinstrum Verification Officers
  3.1 RVO Support section in Rinstrum website
  3.2 Common mistakes
4. Rinstrum Updates
  3.1 UKCA Approvals
  3.2 Rinstrum Onboard Weighing
5. Conclusion

## Slide 3
EU-Regulation
EU- and international Norm
2014/31/EU (NAWI)
2014/32/EU (MID)
EN45501
OIML R60 & R76
WELMEC
Introduction
1.1 Guidelines and Standards
Worldwide harmonization of legal metrology
  OIML recommendations
  OIML documents
  OIML Certificate System
Western European Legal Metrology Cooperation
  WELMEC guides: harmonization of interpretations
  Corresponding tables
EU Regulations
  EU Directives
  Publication of normative documents and harmonized standards

## Slide 4
Terminilogy
scale = measuring instrument = weighing instrument = non-automatic weighing instrument (NAWI)
Definition mentioned in the 2014/31/EU:
Weighing instrument’ means a measuring instrument serving to determine the mass of a body by using the action of gravity on that body. A weighing instrument may also serve to determine other mass-related magnitudes, quantities, parameters or characteristics
‘Non-automatic weighing instrument’ or ‘instrument’ means a weighing instrument requiring the intervention of an operator during weighing;
Introduction
1.2 EU Directives

## Slide 5
Terminilogy
Verfication
Checking if the instrument fullfils all the requirements.
Calibration
Determining how much the measured value differ from the expected value.
Adjustment
Making the error smallest as possible.
Conformity Assestment
Checking the instrument against the requirements in EU directives.
Introduction
1.2 EU Directives

## Slide 6
Measuring instruments (MID)
2014/32/EU
Non-automatic weighing instruments (NAWI)
2014/31/EU
CE Marking
CE Marking can only be placed on an instrument if it complies with EU Directives.
M22 0122
1.2 EU Directive
Introduction

## Slide 7
• Instruments placed on the market must comply with the directives.
• Instruments put in use must have had a conformity assessment.
• At the time of putting into use a declaration of conformity must be available.
1.2 EU Directive
Introduction

## Slide 8
By declaring the conformity of a scale with the essential requirements of the relevant EU regulations, especially 2014/31/EU, the newly placed scale is considered equal to a calibrated scale.
The processed is called
Manufacturer Conformity Assessment process
RINSTRUM is the manufacturer of the scale (indicator and bridge)
The MCA process is closed by RINSTRUM, not by the RVO on site
Therefore, the documents have to be sent to RINSTRUM.
Introduction
1.2 EU Directive

## Slide 9
Read and understand Rinstrum MCA Requirements and Policy (X00A-873)
Valid Commitment as Rinstrum Verification Officer (X00A-812)
Proof of liability insurance (company of the authorized person). Companies need to submit the liability insurance to Rinstrum.
Participation of authorized persons in internal trainings on the manufacturer's conformity assessment procedure
Successful passing of the online test
Well known and at hand: - All relevant EU-regulations (2014/31/EU) and national law - All relevant Norms EN 45501:2015, OIML and WELMEC
1.3 Prerequisites
Introduction

## Slide 10
Rinstrum Documents Related to RVO
  Work instruction - Initial Verification Procedure (X00A-803-170)
  Work instruction - Verification Test (X00A-804-280)
  Work instruction - Flowchart (X008-809-161)
  Template - Verification of Compatibility (X00A-805-120)
  Template - Test Report (X00A-822-269)
  Template - Check List for the Execution of First Verification (X00A-808-180)
The templates used for all documents must be up-to-date, it is your responsibility. Updates will be notified via email.
Please log on to website and get the latest revisions.
1.3 Prerequisites
Introduction

## Slide 11
Other Documents
  Type Approval Certificate (TAC) for the used load cell OIML R60 certified
  Calibration Certificate for the used weights (documenting the traceability to national standards)
1.3 Prerequisites
Introduction

## Slide 12
Training Rinstrum Verification Officer RVO:
RVO will sign a Commitment (X00A-812, X00A-820)
Training Certificate (X00A-811, X00A-819)
Renew authorization once in every 2 year
Rinstrum Verification Officer RVO:
MCA Procedure X00A-803
Sends files to verification@rinstrum.com
MCA Accounting:
Raise the sales order
Acknowledgement email
Rinstrum Metrological Services Assistant:
Store files with the S/N of the indicator in M-Files.
Conformity Check X00A-807
MCA procedure – Internal X00A-860
Warrants the issue of a Declaration of Conformity
Rinstrum Metrological Services:
Declaration of Conformity issued?
Rinstrum Metrological Services Assistant:
Inform RVO on Issues
Raise a Case in M-files for the MCA.
RVO:
Clear the Issue with MCA
Rinstrum Metrological Services Assistant:
Update the Case in M-files for the MCA
Inform RVO.
Rinstrum Metrological Services Assistant:
Issue Declaration of Conformity
Notifies the MCA Accounting.
MCA Accounting:
Is the Company ON Hold ?
MCA Accounting:
Sends the issued Declaration of Conformity with the invoice
RVO:
Sends or gives in person the Declaration of Conformity to the customer. If necessary, according to national law, he will advice the customer to inform the National Weights and Measures Office about the verified scale.
MCA Accounting:
Is the company released ?
MCA Accounting:
Reject the MCA
Raise a Case and inform RVO.
RVO:
The verification isn’t valid and all stickers, sealing stickers and labels must be removed.
1.4 MCA Flow Chart (X00A-809)
Introduction
Yes
Yes
Yes
No
No
No
Yes

## Slide 13
1.5 Rinstrum MCA Team
Introduction
    Marek Swierzy
    General Manager
    Rinstrum Europe GmbH
    Marek.Swierzy@rinstrum.com
    Metrological Service Manager
    Kasun Rathnayake
    Technical Assistant
    Rinstrum Pvt Ltd
    Kasun.Rathnayaka@rinstrum.com
    Metrological Service Assistant
    Sachini Hithakshika
    Sale and Administration Support
    Rinstrum Pvt Ltd Sachini.Hithakshika@rinstrum.com
    MCA Accounting
    Thahiru Herath
    Engineering Manager
    Rinstrum Pvt Ltd
    Thahiru.Herath@rinstrum.com
    Metrological Service Assistant
    Kalana Lalitharathne
    General Manager
    Rinstrum Pvt Ltd
    Kalana.Lalitharathne@rinstrum.com
    General Manager
    Sarala Rathnayake
    Product Development Engineer
    Rinstrum Pvt Ltd
    Sarala.Rathnayake@rinstrum.com
    Metrological Service Assistant

## Slide 14
MCA Procedures and Instructions
2.1 RVO MCA Procedure
 | WI | 
Check list |  | X00A-808
Visual inspection | X00A-803-170 | 
Printer | The printout should be checked | 
Verification of Compatibility for Modules |  | X00A-805
Verification tests | X00A-804-280 | X00A-822
Sealing |  | 
Rinstrum Certification and Trade Label |  | 
Email |  | 

## Slide 15
MCA Procedures and Instructions
Check list X00A-808
2.2 Check List for the Execution of MCA (X00A-808)

## Slide 16
2.3 Verification of Compatibility (X00A-805)
Ensure that the modules are compatible
MCA Procedures and Instructions

## Slide 17
2.3 Verification of Compatibility
MCA Procedures and Instructions

## Slide 18
2.3 Verification of Compatibility
MCA Procedures and Instructions

## Slide 19
2.3 Verification of Compatibility
MCA Procedures and Instructions

## Slide 20
2.3 Verification of Compatibility
MCA Procedures and Instructions

## Slide 21
The tests should be conducted in accordance with DIN EN 45501-2015
The results shall be recorded in the Test Report X00A-822.
At the beginning of each test ensure that the scale is in its reference position (not tilted), i.e. the bubble is in the center of the marked circle.
The tests should be performed at a steady ambient temperature.
At the end of the tests for Test Report, HI_RES should be switched off before the calibration counter is recorded.
2.4 Verification Tests (X00A-804)
MCA Procedures and Instructions

## Slide 22
Visually inspect the scale – indicator , weighbridge
  Damage or other problems which may make the scale unsuitable
  Metrological characteristics, i.e. accuracy class, Min, Max, e, d
  Ensure the settings in the indicator – according to OIML (USE: OIML, zero range and zero band not more than 4% Max)
  Identification of firmware
  Ensure the correct modules are used
  Prescribed inscriptions and positions for verification and control marks
  Ensure that the cable shield and earth are connected. If RS232-interface is used, please advise the manufacturer to install a module to avoid damage because of potential differences.
  Check the construction of the weighbridge to be in accordance with EN45501:2015 and WELMEC 2.4
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 23
Dual Range max. < 1t
Dual Range max. > 1t
Dual Range Truck Scale
Single Range max. < 10kg
Single Range max. < 1t Gravity
Single Range max. > 1t Gravity
Single Range Truck Gravity
Hanging scale < 1t
Single Range medical bed
2.4 Verification Tests
X00A-822 Select the appropriate sheet regarding to the tested scale. At the top of each sheet the language English or German can be chosen. All orange data fields must be completed.
MCA Procedures and Instructions

## Slide 24
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 25
Substitution of standard weights at verification (for Max.>1t)
When testing instruments at the place of use (application), instead of standard weights any other constant load may be used, provided that standard weights of at least 50% Max are used.
If it isn’t allowed to transport more than 20t of weights on the roads, like in Hungary, it is permitted to do Repeatability Test for a scale with Max up to 60t with a contingent of only 20t standard weights. Provided that the repeatability max.error is not more than 0,2e. Repeatability Test must then be performed at a total load of about 50% Max, made up of the standard weights + Truck + substitution material.
If the repeatability error is not more than 0,3e the portion of standard weights may be reduced to 35% Max
If the repeatability error is not more than 0,2e the portion of standard weights may be reduced to 20% Max
Procedure for using substitution material:
Put the contingent of standard weights onto the scale that you would like to substitute. Note down which weight the indicator shows. Then load substitution material onto the scale until the indicator shows exactly the same weight.
After that, gently add the standard weights up to the next test load. Same procedure would be downwards.
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 26
Tests
Repeatability Test
Accuracy of zero device
Accuracy of tare device
Weighing / Linearity Test
Tare (Weighing Test)*
Eccentricity Test
Test with coasting load (Truck Scales)
Tilting test ( for medical bed Class III only )
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 27
Test 1: Repeatability Test
The instrument’s automatic zero-setting and zero-tracking device may be in operation during this test.
Before testing repeatability the scale has to be pre-loaded with Max.
At least three weighings with about 80% Max are necessary.
It is allowed to zero the instrument between the weighings, if it doesn’t show zero.
Evaluation of Error is calculated in the sheet. The error between the three readings shouldn’t be more than the absolute value of the permissible error.
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 28
2.4 Verification Tests
Actual sample load
Indication
Indication Error
Max Permissible Error
MCA Procedures and Instructions

## Slide 29
Test 2: Accuracy of zero device
High-Res Mode must be switched off.
A weight close to Zero must be loaded onto the scale.
Indicator must now be set to Zero by pressing the Zero-button
A weight close to Zero must be loaded onto the scale.
Note down the number of added weights.
Evaluation of Error is calculated in the sheet.
MCA Procedures and Instructions
2.4 Verification Tests

## Slide 30
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 31
Test 3: Accuracy of tare device
High-Res Mode and Zero tracking must be switched off.
This test is to be performed in the same manner as Test 2.
A weight between 1/3 and 2/3 of Max must be loaded onto the scale.
Indicator must now be tared by pressing the Tare-button.
Standard weights of 1/10e must be gently added, until indication changes from one scale interval to the next above.
Note down the number of added weights.
Evaluation of Error is calculated in the sheet.
MCA Procedures and Instructions
2.4 Verification Tests

## Slide 32
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 33
Test 4: Weighing / Linearity Test
Apply 5 test loads and similarly remove them back to zero.
The loads selected shall include Min and Max and values at or near to those at which the maximum permissible error (mpe) changes (500e and 2000e).
The scale should not be allowed to return to zero between loads.
If substitution material is used, you must first load from zero up to the maximum quantity of standard weights. Then remove the weights and substitute the previous weights with substitution material. Repeat the procedure until Max is reached. Unload in reverse order to zero.
The calculation sheets calculates the total error Eerror - Ezero. If this total error is less or equal to 0,5e, Tare Test 5 has not to be performed.
MCA Procedures and Instructions
2.4 Verification Tests

## Slide 34
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 35
Test 5: Tare (Weighing Test)
The Test Report sheet calculates if this Test has to be undertaken or not.
This test has to be performed if the total error in Test 4 is more than 0,5e.
A load between 1/3 and 2/3 Max should be applied and then be tared.
Apply 5 test loads and similarly remove them back to zero.
The loads selected shall include Max (maximum possible NET load) and Min and values at or near to those at which the maximum permissible error (mpe) changes (500e and 2000e).
The scale should not be allowed to return to zero between loads.
If substitution material is used, you must first load from zero up to the maximum quantity of standard weights. Then remove the weights and substitute the previous weights with substitution material. Repeat the procedure until Max is reached. Unload in reverse order to zero.
MCA Procedures and Instructions
2.4 Verification Tests

## Slide 36
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 37
Test 6: Eccentricity Test
Zero-setting or zero-tracking device shall not be in operation.
Large weights should in preference to several small weights.
For an instrument with a load receptor having not more than four points of support, the four quarter segments roughly equal to ¼ of the surface of it shall be loaded in turn. The test load shall be 1/3 Max.
For an instrument with a load receptor having more than four points of support, the test load of (1/n-1) Max shall be applied over each support on an area of 1/n of the surface area of the load receptor.
Where two points of support are too close together for the test load, the doubled load shall be distributed over twice the area on both sides of the axis connecting the two points of support.
Eccentricity for Dual range scale should be tested at 1/3 of Max.
MCA Procedures and Instructions
2.4 Verification Tests

## Slide 38
MCA Procedures and Instructions
2.4 Verification Tests

## Slide 39
MCA Procedures and Instructions
2.4 Verification Tests

## Slide 40
Test 7: Test with coasting load (Truck Scales)
A truck load of at least 50% Max but not exceeding 80% Max shall be driven onto the weighbridge from each of the access roads.
Measurements shall be carried out in each of the three bridge segments.
The distance from the middle position to the first or end position shall be more than 0,1B , where B is the length of the bridge.
Length an width of scale should be entered.
MCA Procedures and Instructions
2.4 Verification Tests

## Slide 41
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 42
Test 8: Tilting test ( for medical bed Class III only )
The bed must be in reference position, without inclination, no load on it.
Indicator must now be set to Zero by pressing the Zero-button.
Max. load must be put onto the bed and the indicated weight must be written into the test report.
Then, the bed must be inclined until the bubble in the levelling device touches the marking in all four directions.
Now the indicated weight without any load and with Max load must be written to the test report in every inclined direction.
This test is not necessary for medical beds Class IIII
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 43
3.4 Verification Tests
MCA Procedures and Instructions

## Slide 44
Earth Gravity
If necessary, verify the earth gravity for the place of use
Place of installation
Name and town of the installation
Calibration Counter
The high resolution mode (HI_RES) shall be switched off. Then calibration counter should be recorded.
2.4 Verification Tests
MCA Procedures and Instructions

## Slide 45
Trade Label RVO 2023
X00A-309-101 (A11030-23)
Certification Label
X00A-302-202 (A11047)
Certification Seal
X00A-305-101 (A11029)
Transparent Foil
X00A-306-100 (A11037)
Use this Stockcode LABEL-23-KIT for set of labels for 10 Scale.
  50 x Certification Seals
  10 x Trade Labels RVO 2023
  10 x Certification Labels
  20 x Transparent Foils
Manufacture Conformity Assessment
2.5 Seals and Labels

## Slide 46
How to put label and seal on the indicator
2.5 Seals and Labels
Manufacture Conformity Assessment
Affix the verification labels. Ensure they are filled out correctly.
Use a UV rated permanent pen.
Stick the transparent foil over the verification labels for protection.
Seal all the affixed labels.

## Slide 47
Sealing the R320 Indicator
2.5 Seals and Labels
Manufacture Conformity Assessment
Aflix Trade lable and certification Lable with void labels
Option 1
Option 2
Void Lable
Void Lable
Void Lable
Option 3
Void Lable
Void Lable
Void Lable

## Slide 48
2.5 Seals and Labels
Manufacture Conformity Assessment

## Slide 49
Sealing the R420 Indicator
2.5 Seals and Labels
Manufacture Conformity Assessment
Aflix Trade lable and certification Lable with void labels
Option 1
Option 2
Void Lable
Void Lable
Void Lable
Void Lable
Void Lable
Option 3
Void Lable
Void Lable
Option 4

## Slide 50
2.5 Seals and Labels
Manufacture Conformity Assessment

## Slide 51
Sealing the Junction Box
2.5 Seals and Labels
Manufacture Conformity Assessment

## Slide 52
The filenames of X00A-805,808,822 must be: SheetNumber_Series of Indicator_SerialNumber_VerificationOfficer e.g. “X00A-816_R320-K356_3131313_rvo12121.pdf”.
Email to verification@rinstrum.com with the following details in email subject: “Series of Indicator_Indicator-SerialNumber_SerialNumber of Scale_RVO-No.”
The data must be unzipped
Only data of one scale must be attached to one email
2.6 Documents to Rinstrum
Document | Format
X00A-808 | Pdf or scanned Pdf\Image
X00A-805 | Excel or Pdf
X00A-822 | Excel or Pdf
Photos | Images
Certificate of used standard weights (Copy) | Pdf
Certificate of the used delta-weights (Copy) | Pdf
TC Loadcell (Copy) | Pdf

## Slide 53
Rinstrum will issue the Declaration of Conformity after processing.
Send the Declaration of Conformity to the customer.
According to national law, if requested, RVO should ask the customer for informing the local authority about the verified scale and sending all requested documents.
2.7 Declaration

## Slide 54
2.8 Clarification and Rejection
When there is an issue with the MCA files, it will be on clarification state and RVO will be informed regarding this.
This will extend the processing time of the MCA.
Once RVO clears the issue, MCA will be processed, and Declaration of conformity will be issued.
If the issue can’t be resolved, MCA will be rejected. This will be informed to RVO and stickers places in the scale should be removed.
RVO can perform a new MCA for the scale and send documents to the verification email.

## Slide 55
3.1 RVO Support section
Rinstrum Verification Officers
RVO Web Page is accessible through https://rinstrum.com/product-category/rvo/

## Slide 56
3.1 RVO Support section
Rinstrum Verification Officers
Certificates
MCA Procedure
RVO News
RVO Newsletter
Training
Calibration
R400 Calibration procedure
R300 Calibration procedure
All related MCA Process Certificates
MCA Process document templates and work instructions.
Latest news including MCA updates and Upcoming events
Monthly Newsletter Archive
Training Material

## Slide 57
3.2 Common Mistakes
Rinstrum Verification Officers
Incorrect Type Approval Certificate Number
Rinstrum List of certificates
R300 V4 Indicator vs R300 V4 HW1.2 Indicator
 | T Number | TC Number
R300 Series V4 | T6244 | TC6242
R300 V4 HW 1.2 | T11438 | TC11437
X300 Series | T7418 | TC6821
R400 Series | T6822 | TC6821
C500 Series | T11766 | TC11766

## Slide 58
3.2 Common Mistakes
Rinstrum Verification Officers
Imperfect Sealing
Make sure to seal in a way that Scale can not be
Modified without removing the void stickers.
Send all images with the MCA file.

## Slide 59
4.1 UKCA Approvals
Rinstrum Updates
Rinstrum has applied for UKCA approvals and currently under processing.
RVOs can apply for a training with Rinstrum for performing MCA with UKCA approvals.
Email us at varification@rinstrum.com for more information.

## Slide 60
4.1 Rinstrum Onboard Weighing Approval
Rinstrum Updates

## Slide 61
4.1 Rinstrum Onboard Weighing Approval
Rinstrum Updates

## Slide 62
4.1 Rinstrum Onboard Weighing Approval
Rinstrum Updates
R420 Indicator
M4907
Tilt Sensor Module
M4211
Tilt Compensation module
M4501
Data Storage Module

## Slide 63
Conclusion
Conclusion
This training is focused on refreshing the knowledge of the RVO to perform successful MCA.
Training includes an online test.
Renew of RVO commitment (X00A-812).
RVO certificate

## Slide 64
X00A-837-201